HPLC Operating Instructions

Turn on the HPLC.  Prime the pumps by drawing the solvent (Pump B)

Use a blunt needle syringe for injections.  Ensure loading device is set to ‘L’ (for load).

Wash syringe a few times with your sample, tap-out/remove any air bubbles in the syringe.

Draw about 2 – 3mL of the ‘test’ sample into the syringe.

Make sure there is sufficient amount of the solvent in the solvent reservoir and the eluate ends up in the waste container.

Running the Program:

Double Click ‘Biologic’ Icon on the Desktop.  Once the program is initiated, on the opening screen make sure to set Gradient Pump settings as A=0 and B=100; flow rate 0.08 mL/min.  SV5-4 @ 1, AVR7-3 @L (L for loop).

Press [START] and [Zero the Base line]. This operation flushes the system and can be stopped after a couple of minutes. Then on, the menu buttons to click etc. are as follows.

‘Browser’ button (upper left corner)

‘USERS’ → ‘Chem 380 Lab’ → Project → ‘Method 1.3’;

(‘Method 1.3’ has already been programmed).

With ‘Method 1.3’ highlighted, click ‘New’ → ‘New Run’

Title (meaningful unique file name) the run at this point;  → ‘OK’;

Inject the sample (into the loop);’ → ’Start’

[Inject the sample into the loop of the injector, through a 45 filter disc.  Please do not force the needle into the injector; you may leave the needle and syringe in the injector. Only a volume equal to that of the loop size (50L) will be injected on to the column at the running of the chromatogram.]

Pay attention to the instruction/caution boxes and the Protocol bar (bottom of screen)

Note:  Analysis length is set to ~15 min. However caffeine elutes ~ 8 min. under the parameters of the experiment. One may click the ‘abort’ button to stop the detector sampling the eluate (stop the run) at a suitable time past complete elution. Once the separation is complete use the vertical scroll bar of the upper left window (section next to the graph) to zoom in/out on the peak.

Click ‘Post Run’ button, a cursor appears on the chromatogram window.

Place the cursor on top of the peak(s).  Left click to place a tag on the top of the peak(s) of the chromatogram. Watch the information on the top left window for the highest absorbance of the peak, before selecting the peak.  At this time you may save the data file; File > Export Data.

Click ‘Tags’ button; UV absorbance at the peak(s) of the eluted components selected in the preceding step appear in a table.  This value of absorbance is proportional to the concentration of the analyte (plug) flowing through the detector.

To use the same method and run a new sample: Load the sample into the loop.  Select ‘File’ → ‘New Run’ ……